## Slide 1
VK-PA-Pico Power Analyzer
Picoammeter Mode - Operating Principle to Measure Short Circuit Current (Isc) of Solar Cell
+
-
Solar Cell
To Analog to Digital Converter
0 V
0 V (Short Circuit condition)
R
Feedback Ammeter
I

## Slide 2
VK-PA-Pico Power Analyzer
MPPT Mode - Operating Principle
ADC :- Analog to Digital Converter DAC :- Digital to Analog Converter
Solar Cell
R
To ADC (for Current Reading)
Feedback Ammeter
I
+
-
Current Sink
Voltage Sense
Current Sink
Controller
To ADC (for Voltage Reading)
Control Signal From DAC
V
I

## Slide 3
VK-PA-Pico Power Analyzer
Maximum Power Point Tracking Mechanism
Output Voltage
Output Power
V-ΔV V V+ΔV
P1
P2
P3
S1-2
S2-3
If (S1-2 > 0 and S2-3 < 0) then
Maximum Power Point is P2
If (S1-2 > 0 and S2-3 > 0) then
P1  P2
P2  P3
Find New P3 at V+2ΔV
If (S1-2 < 0 and S2-3 < 0) then
P3  P2
P2  P1
Find New P1 at V-2ΔV
P1
P2
P3
P1
P2
P3
User can set the desired ΔV value. Default value is 1 mV

## Slide 4
VK-PA-Pico Power Analyzer
VK-PA-Pico can also work as current source if needed. In the potentiostat mode it work as current source to the device (such as resistor or diode) connected to analyzer outputs.